Py-Spike2 Trigger

Sending digital markers from Python to Spike2 through a BBTK USB TTL module and a CED 1401.

This document is oriented towards computers that run on Windows. Certain steps will need to be modified slightly to work within MacOS.

1. The BBTK-TTL driver

The TTL module appears on the experiment computer as a COM port. Windows usually installs the driver by itself, so check before downloading anything.

- Plug the module into the computer and wait for Windows to finish.

- Look at the red and green LEDs on the TTL module when you plug it in, they should both light up for ~ 10s, if they do then you know it is plugged in properly.

2. Open Device Manager: press the Windows key + X, and find Device Manager.

3. Expand Ports (COM & LPT). If you see USB Serial Port (COMn), the driver is already installed. Make note of the number (e.g. COM2)

4. If you can’t find it, look under Other devices or Universal Serial Bus controllers for an item with a warning triangle, usually called FT232R USB UART. This means the driver it needs is missing so carry on to step 5.

5. Go to https://ftdichip.com/drivers/vcp-drivers/

6. Find the row for your version of Windows and download the setup executable, not the plain ZIP.

7. Unplug the module, run the downloaded installer, and accept the defaults.

8. Plug the module back in and repeat step 3. It should now appear as a COM port.

9. Right-click USB Serial Port (COMn), choose Properties, then the Port Settings tab, then Advanced. Set Latency Timer (msec) to 1 and click OK. The default is 16 ms and it delays every single marker by that much.

- This just makes sure that the markers send exactly when you want them to. If a particular amount of latency from marker initiation is pertinent to your study, this is the easiest way to modify that as well.

10. Always plug the module into the same USB socket. Windows gives the COM number per socket, so it will stay the same from session to session.

2. The cables

1. Plug the USB-A end of the USB cable into the experiment computer.

2. Plug the USB-B end into the TTL module.

3. Plug the DB25 female end of the supplied cable onto the DB25 male connector on the TTL module.

4. Plug the DB25 male end into Digital Inputs on the back of the CED 1401.

About the current DB25 cable: it will carry the timing properly but not a specific code, so every marker in Spike2 will read AA. That is normal and nothing is wrong. The marker times are correct, and section 6 explains how to label them afterwards.

There is technically ways to create a special cable that can carry specific hex-bytes (e.g. Aa, Ab, Ac), but that is out of my breadth at this current moment.

3. Spike2 configuration

On the Spike2 computer. Steps 1 to 5 are done once and saved; step 6 is a display setting.

1. Edit, then Preferences, then the Sampling tab. Tick Use rear connector for events. This only needs doing once per computer.

2. Open your sampling configuration: Sample, then Sampling Configuration.

3. On the Channels tab, click New Channel and choose Marker. In some versions of Spike2 it appears in the config. as Digital Marker but one the date of this document (Sept. 25/2026 it is Marker).

4. Set the channel to 32 (a digital marker is always channel 32), the title to Trig and the maximum event rate to 100. The title and maximum event rate are technically up to you, this is just a baseline for standardization purposes.

5. File, then Save Configuration. Then at the start of every session you’ll go to File, Load Configuration and select this file.

6. After the first recording, right-click the Trig channel, choose Draw Mode, and pick a mode that shows the marker code as text so you can see the marks.

4. Python

4.1 The trigger module

import time

BAUD         = 115200   # the speed the BBTK runs at; do not change

CMD_RESET    = b"RR"    # resets the module and clears its output lines

CMD_PING     = b"##"    # the module replies "XX" if it is awake

PING_REPLY   = b"XX"

LINE_UP      = b"80"    # strobe line high: nothing is being sent

LINE_DOWN    = b"00"    # strobe line low: the 1401 records a marker

PULSE_MIN_MS = 5        # how long the line stays down

RESET_WAIT_S = 0.1      # pause after the reset, before anything else

class Trigger:

def __init__(self):

self._ser = None

self.enabled = False   # True once the port is open

self.message = ""      # what happened, for printing on screen

self._down = False     # is the line currently down?

self._down_at = 0.0

self.epoch = time.perf_counter()

def open(self, port):

"""Open the COM port and park the line up."""

try:

import serial

except ImportError:

self.message = "pyserial missing: pip install pyserial"

return False

try:

self._ser = serial.Serial(port, BAUD, timeout=1,

write_timeout=0.05)

self._ser.write(CMD_RESET)

time.sleep(RESET_WAIT_S)

self._ser.reset_input_buffer()

self._ser.write(CMD_PING)

answered = PING_REPLY in self._ser.read(2)

self._ser.write(LINE_UP)

except Exception as exc:

self.message = f"{port} did not open: {exc}"

return False

self.enabled = True

self.epoch = time.perf_counter()

if answered:

self.message = f"BBTK answered on {port}"

else:

self.message = f"{port} opened, but the BBTK did not answer"

return True

def pulse(self):

"""Send one marker. Returns the time in ms, or None."""

if not self.enabled:

return None

if not self._write(LINE_UP):      # make sure the line is up

return None

if not self._write(LINE_DOWN):    # drop it: this is the marker

return None

self._down = True

self._down_at = time.perf_counter()

return round((self._down_at - self.epoch) * 1000.0, 1)

def service(self):

"""Call once a frame. Puts the line back up after a marker."""

if not self._down:

return

gap = (time.perf_counter() - self._down_at) * 1000.0

if gap >= PULSE_MIN_MS:

self._write(LINE_UP)

self._down = False

def close(self):

"""Park the line up and release the port. Safe to call twice."""

if self._ser is not None:

try:

self._ser.write(LINE_UP)

self._ser.close()

except Exception:

pass

self._ser = None

self.enabled = False

def _write(self, data):

"""All writes go through here so a failure cannot crash."""

try:

self._ser.write(data)

return True

except Exception as exc:

self.enabled = False

self.message = f"trigger link lost: {exc}"

return False

What each part does:

- The constants at the top. The BBTK is controlled by two-character commands sent down the COM port. "80" puts the strobe line up, "00" puts it down. The 1401 records a marker the moment the line goes down.

- open(port). Opens the COM port, resets the module, asks it whether it is there, and parks the line up. It returns True or False and puts a plain-English result in trig.message instead of crashing. So if you get the COM # wrong for whatever reason, this keeps your session running without it.

- pulse(). Sends one marker, as two quick writes: the first makes sure the line is up, the second drops it. The drop is what Spike2 records. It returns the time in milliseconds since the port was opened.

- service(). Puts the line back up again a few milliseconds after a marker. This must be called every frame, because if the line never goes back up there is no drop for the next marker and then you can’t send more markers.

- close(). Parks the line and releases the port. Safe to call more than once, so it can go into like an end block.

- _write(). Every write goes through this one function, so if the cable is pulled out mid-session the link simply switches off and records why, rather than raising an error into your experiment.

This should be its own .py file, separate from your main experiment file. As long as both files are present in the same folder there should be no issues. For the sake of example purposes, let’s call the above code trigger.py.

Check it works

1. Start sampling in Spike2.

2. Run your Python script, or a few test markers.

3. Marks should appear on the Trig channel at the right times, each reading AA.

4. If nothing appears, check in this order: the DB25 is in Digital Inputs and not another socket, the module is powered and its lights are on, and the COM number in your Python script matches Device Manager.

4.2 Setting it up in your study

Near the top of your script, once:

from trigger import Trigger

trig = Trigger()

trig.open("COMn")     # the number from Device Manager

print(trig.message)       # this should say the BBTK answered

In every loop that draws frames:

while running:

draw_everything()

pygame.display.flip()

trig.service()        # puts the line back up after a marker

And at the end, on every way out of the script including errors and the quit key:

trig.close()

4.3 Sending a marker

The rule is the same everywhere: trig.pulse() goes on the line immediately after whatever makes the event happen, never before it.

After something appears on screen:

screen.blit(stimulus, rect)

pygame.display.flip()        # the participant can now see it

t = trig.pulse()             # marker

On a keypress or a response:

for event in pygame.event.get():

if event.type == pygame.KEYDOWN:

t = trig.pulse()     # marker first, before anything else

handle_response(event.key)

At an event with no screen change, such as a sound or the start of a block:

bell.play()

t = trig.pulse()

5. PsychoPy

Exactly the same module and the same three calls. Keep trigger.py in the same folder as your PsychoPy experiment.

Coder

from trigger import Trigger

trig = Trigger()

trig.open("COM12")

print(trig.message)

stim.draw()

win.flip()          # the participant can now see it

t = trig.pulse()    # marker

trig.service()      # once every frame

trig.close()        # at the end

Builder

Add a Code Component to your routine and put these lines in these tabs:

| Tab | What goes in it |
| --- | --- |
| Begin Experiment | from trigger import Trigger trig = Trigger() trig.open("COMn") print(trig.message) |
| Begin Routine | t = trig.pulse() |
| Each Frame | trig.service() |
| End Routine | thisExp.addData("marker_ms", t) |
| End Experiment | trig.close() |

- Put trig.service() in the Each Frame tab of every routine that can be running when a marker is sent.

- Put the Code Component above your stimulus components in the routine, so its code runs first.

- Builder has no setup screen, so read the console after print(trig.message) before you bring the participant in.

6. Labelling the markers in Spike2

Every mark on the Trig channel reads AA, so you tell them apart by their position in the sequence: they arrive in a fixed order at known intervals. A short Spike2 script turns them into readable labels after the session.

What the script needs to do:

1. Read every marker time from the Trig channel into an array.

2. Work out which marker is which from its position. If your experiment sends three markers per trial, markers 1 to 3 are trial 1, markers 4 to 6 are trial 2, and so on.

3. Check the gaps between markers against what your experiment actually sends. This matters: if one marker is missed and you only count positions, every label after it is wrong.

4. Create a new TextMark channel to hold the labels.

5. Write one label at each marker time, for example 1a, 1b, 1c, 2a and so on.

6. Print a summary saying how many markers were found and how many were expected, so a bad recording is obvious straight away.

Skeleton:

var ch%, n%, t[10000], i%, lab%;

ch%  := 32;                               'the Trig channel

n%   := ChanData(ch%, t[], 0, MaxTime()); 'all the marker times

lab% := MemChan(8);                       'a new TextMark channel

ChanTitle$(lab%, "Labels");

for i% := 0 to n%-1 do

'three markers per trial: 1a 1b 1c, 2a 2b 2c, ...

MemSetItem(lab%, 0, t[i%], 0,

Str$(i%/3 + 1) + Chr$(97 + i% mod 3));

next;

ChanShow(lab%);

PrintLog("markers found: %d\n", n%);

Reading Spike2 script

Spike2 has its own script language. It looks a bit weird, but the skeleton above only uses a handful of rules:

Names end in a symbol that says what they hold. % means a whole number (n%, i%). $ means text (ChanTitle$, Str$). No symbol means a decimal number, like the times in t[], which are in seconds.

:= stores a value. ch% := 32 means “ch% is now 32”. A plain = only compares two values.

; ends every statement, a bit like a full stop.

' starts a comment. Spike2 ignores the rest of that line, so anything after ' is a note for people.

[ ] means an array, a numbered list of values. t[10000] makes room for 10,000 values; t[i%] is the value at position i%. Counting starts at 0, so the first value is t[0].

for ... do and next; make a loop. Everything between them runs once for each value of the counter.

var declares variables. Every variable must be listed with var before it is used.

The skeleton EXPLAINED  (because actually what)

var ch%, n%, t[10000], i%, lab%;

This is essentially defining the variables that the script is going to use. ch% is the channel to read, n% will hold how many markers were found, t[10000] is room for up to 10,000 marker times (you can make this whatever but 10,000 felt safe), i% is the loop counter and lab% will hold the number of the new label channel.

ch%  := 32;

The Trig channel. A digital marker channel is always channel 32, so this only changes if you recorded the markers somewhere else.

n%   := ChanData(ch%, t[], 0, MaxTime());

Copies the time of every marker on channel 32 into the array t[], from 0 seconds to the end of the file (MaxTime()). You can change MaxTime if you want a certain script for a certain time intervale. It returns how many it copied, which is stored in n%. Afterwards t[0] is the first marker and t[n%-1] the last. Writing t[] with empty brackets passes the whole array, not one value.

lab% := MemChan(8);

Creates a new, empty channel. 8 is the channel type TextMark: a marker that can carry a short piece of text. It returns the new channel’s number, stored in lab%.

ChanTitle$(lab%, "Labels");

Names the new channel “Labels” so you can find it on screen.

for i% := 0 to n%-1 do

Goes through the markers one at a time. i% starts at 0 (the first marker) and stops at n%-1 (the last), because counting starts at 0.

MemSetItem(lab%, 0, t[i%], 0,

Str$(i%/3 + 1) + Chr$(97 + i% mod 3));

Adds one label to the Labels channel. The parts, in order: lab% is the channel to write to; 0 means add a new item rather than replace one; t[i%] is the time, so the label sits exactly on the marker; the next 0 is a marker code, which is not used here; the last part is the text.

The text is built from the marker’s position. i%/3 divides by 3 and, because both are whole numbers, throws away the remainder: markers 0, 1, 2 give 0, markers 3, 4, 5 give 1, and so on. Adding 1 gives the trial number, and Str$ turns it into text. i% mod 3 is the remainder, which cycles 0, 1, 2. Chr$ turns a character code into a letter, and 97 is the code for “a”, so Chr$(97), Chr$(98), Chr$(99) are a, b, c. The + joins the two pieces of text.

Example: the fifth marker has i% = 4. 4/3 is 1, plus 1 is 2. 4 mod 3 is 1, and Chr$(97 + 1) is “b”. Its label is “2b”.

next;

The end of the loop. Spike2 goes back to the for line with the next value of i%.

ChanShow(lab%);

Makes the new Labels channel visible in the data window.

PrintLog("markers found: %d\n", n%);

Writes a line to the Log window (View, then Log). %d is replaced by the value of n%, and \n starts a new line.

Save it with a .s2s extension and run it with Script, then Run. Check the exact arguments against Spike2’s own script help (Help, then Script Language) before relying on it, as they vary slightly between versions.

You can run the script at the very end of all of your data collection. Say for example you have 40 .smrx files containing EMG data. As long as the trigger logic is consistent across those files, and the script expects that, you can run the labelling script for the entire folder at once.